TP60 instructions

Liquid nitrogen orders

- When the TP60 level arrive between the two transparent lines just above the red line, it is time to order the other dewar refilling. When the TP60 level reaches the red line, it is time to exchange the dewar.

- The instructions to order liquid nitrogen « Liquid nitrogen order » are on the board next to the entrance to the control room. You have to order the « pequeño dewar »

Preparation of a full TP60

When you have received the refilled TP60, proceed as follows:

- Put on PPE (Gloves/Goggles/Corresponding clothes)

- Verify that transfer’s valves are positioned as shown on the next picture

Transfer listo para ser montado en el dewar

- Insert the green seal on the transfer and hold it at the top to be sure it does not come into contact with the liquid nitrogen

- SLOWLY go down the transfer into the dewar

- Position the manometer in front of the yellow sticker « position manomètre »

- Place the collar in the same position as the empty dewar and tighten it

- Closet valve N°3

- Open valve N°4 in order to pressurize the dewar

The dewar must be in this configuration:

TP60 ready to be exchanged

Exchange of the empty TP60 by a full one

- Put on PPE (Gloves/Goggles/Corresponding clothes)

- Close valves N°1 and N°4

- Loosen the brass nut

- Remove the white strap

- Take the empty TP60 out of its housing

- Place the full TP60 so that the N°1 outlet is in front of the transfer hose

- Tighten the white strap

- Screw the brass nut

- Open valves N°1 (and N°4 if it’s not done yet)

The dewar must be in this configuration:

Full dewar ready and in its housing

Disassembly of the empty TP60

- Put on PPE (Gloves/Goggles/Corresponding clothes)

- Close the valve N°4 (if it’s not done yet)

- Make sure that there is nothing within 2m of valve N°3 and then open it to release pressure

- Let the pressure drop to 0 bar

- Loosen the collar, remove it and place it on the cardboard plate marked TP60

- GENTLY pull out the dewar’s transfer and place it on the cardboard plate marked TP60

- Remove the green seal and place it on the cardboard plate marked TP60

- Put the cap on the dewar

The empty dewar is ready for refilling

(WHEN THE TRANSFER IS NOT MOUNTED ON THE DEWAR, THE VALVE N°4 MUST ALWAYS BE CLOSED)

« Manual » refill with TP60

ONLY IN CASE OF EMERGENCY OR TIMER FAILURE

If for some reason Coralie gets warmer, it should be manually refilled

Timer for manually refill

- Follow the procedure displayed against the wall next to the timer

Procedure for manual refill